## GC Checklist
SMITH ANALYTICAL, LLC
PO BOX 278, SEABROOK, TEXAS 77586
281-667-6463
WWW.SMITHANALYTICAL.COM
Purpose of this document is to provide a checklist for Analyzer Professionals to insure all proper steps are taken BEFORE the Analyzer is calibrated. In the majority of the cases (>95%) where Statistical Quality Control Chart (SQC) are used and analyzers are shown to be "In Control", when the analyzer goes "Out of Control" something has changed and this needs to be addressed before calibration of the unit is allowed. Additional items can and should be added for your particular location. ALL ITEMS MUST BE REVIEWED BEFORE RECALIBRATING AN ANALYZER THAT WAS PREVIOUSLY IN CONTROL.
SQC "OUT OF CONTROL" CHECK LIST
ITEM | DESCRIPTION | OK - YES/NO | N/A | Comment
1 | Was the Analyzer "In Control" Before This Calibration? |  |  | If Yes, then you have a New Problem. If No, then NO assumptions regarding system performance should be made.
2 | Is the valve gas regulator set per the specifications?
3 | Is the valve gas regulator pressure stable during switching operations? |  |  | If No, this needs to be corrected.
4 | Oven air regulator properly set?
5 | Main Carrier Gas Regulators properly set? |  |  | If No, this needs to be corrected and the validation test ran again.
6 | Carrier Gas Regulator on the GC properly set?
7 | Calibration Gas Regulator on the GC properly set?
8 | Is the calibration bottle pressure below 200 PSIG? |  |  | If Yes, you should consider changing the bottle and then running the validation test again.
9 | Sample System operating at the designed parameters of temperature, flow and pressure? |  |  | If No, correct this and run the validation test again.
10 | Have all standard operating parameters been reviewed to insure everything is correct? |  |  | If No, this needs to be done BEFORE proceeding. Make NO assumptions. Check the Factory documentation to insure everything is correctly set.
11 | Has anyone worked on the gas chromatograph recently? |  |  | If Yes, you need to understand everything this individual did to the unit. They possibly introduced a problem.
12 | Was the Carrier Gas Changed?
13 | Is the baseline noisey? |  |  | If Yes, then its possible you have dirty carrier or bad air (air treator). Also possible you have a sample valve leaking.
14 | Was the Calibration Gas Bottle Changed?
15 | Has the Calibration Bottle Expired? |  |  | If Yes, organizational changes or procedures need to be put into place to prevent this from happeing again.
16 | Is the analyzer repeating on calibration gas? |  |  | Here the interest is precision of the analysis even though the unit is out of control. If the analyzer is NOT repeating on calibration gas, you have a "Y" axis issue.
17 | Is the calibration sample being introduced to the analyzer in the same manner as done previously? |  |  | Answer should be Yes. If No, need to understand why?
18 | Is this analyzer subjected to a multi-point calibration / validation? |  |  | If Yes, how does the unit perform on the other calibration standards? If it is in control on the other standards, the problem is related to the calibration gas being used for the SQC test.
19 | Has the Calibration Bottle Been "Rolled"? |  |  | Some calibration bottles (H2S, Acetylene) tend to stratify when they sit in the shade or might be exposed to cold weather for periods > 24 hours. The bottle needs to be at room temperature and stirred or rolled to insure a homogenous mixture.
20 | Is the Oven Temperature Stable?
21 | Is the Carrier Flow Stable and Per Specification?
22 | Has peak rention time changed since the last analysis?
23 | Are peak retention times changing from cycle to cycle? |  |  | If Yes, you have an "X" axis issue.
24 | Is the sample valve operating properly?
25 | Are the analyzer flows (detector vents) vented to atmosphere? |  |  | If No, then there might be a problem with backpressuring the system.
26 | For gas phase samples, have you tested or checked you Sample Shut-off Valve and Atmospheric Reference Valve? |  |  | If the reading in question is above the Uper Control Limit, this should be considered as a possible cause.
27 | Only one peak is out of control? |  |  | If Yes, then its possible a valve time (backflush or heartcut) is set incorrectly. If more than one component is out of control, there is a larger issue at hand.
28 | Check Backflush Time to insure it is correct?
29 | Check Heartcut Time to insure it is correct? |  |  | You may have to use a "Spotter Blend" to perform this step.
30 | Does the GC have a Thermo-Conductivity Detector? |  |  | If Yes, perform a bridge test.
31 | For Flame Ionization or Flame Photometric Detectors, did you check the flows to insure they are properly set? |  |  | Pressure indication is NOT the sole consideration here. Need to insure flows are properly set.
32 | Does the GC have a Flame Ionization Detector? |  |  | If Yes, perform a visual inspection test of the detector to see if there is any carbon build up. Perform FID electonics test.
33 | Does the GC have a Flame Photometric Detector? |  |  | If Yes, check you air and hydrogen flows to insure they are at the correct ratio. Call Factory or Specialist for help.
34 | Did you correct or discuss this problem with the Plant Analyzer Specialist BEFORE proceeding with recalibration of the system?
REVISION 0 - OCTOBER 2012
©Smith Analytical, LLC. Can only be used by specified Clients of Smith Analytical. Use by others without permission from Smith Analytical, LLC or for Commerical Applications without prior consent is strictly prohibited.

## CEMS
SMITH ANALYTICAL, LLC
PO BOX 278, SEABROOK, TEXAS 77586
281-667-6463
WWW.SMITHANALYTICAL.COM
Purpose of this document is to provide a checklist for Analyzer Professionals to insure all proper steps are taken BEFORE the Analyzer is calibrated. In the majority of the cases (>95%) where Statistical Quality Control Chart (SQC) are used and analyzers are shown to be "In Control", when the analyzer goes "Out of Control" something has changed and this needs to be addressed before calibration of the unit is allowed. Additional items can and should be added for your particular location. ALL ITEMS MUST BE REVIEWED BEFORE RECALIBRATING AN ANALYZER THAT WAS PREVIOUSLY IN CONTROL.
SQC "OUT OF CONTROL" CHECK LIST
ITEM | DESCRIPTION | OK - YES/NO | N/A | Comment
1 | Was the Analyzer "In Control" Before This Calibration? |  |  | If Yes, then you have a New Problem. If No, then NO assumptions regarding system performance should be made.
2 | Have all standard operating parameters been reviewed to insure everything is correct? |  |  | If No, this needs to be done BEFORE proceeding. Make NO assumptions. Check the Factory documentation to insure everything is correctly set.
3 | Are all switches in the proper setting for the analyzers, flow drawer, etc.? |  |  | Check range switches and flow drawer switches.
4 | Has anyone worked on the CEMS recently? |  |  | If Yes, you need to understand everything this individual did to the unit. They possibly introduced a problem.
5 | Was the Calibration Gas Bottle Changed?
6 | Is the calibration bottle pressure below 200 PSIG? |  |  | If Yes, you should consider changing the bottle and then running the validation test again.
7 | Is the analyzer repeating on calibration gas? |  |  | Here the interest is precision of the analysis even though the unit is out of control. If the analyzer is NOT repeating on calibration gas, you have a "Y" axis issue.
8 | Is the calibration sample being introduced to the analyzer in the same manner as done previously? |  |  | Answer should be Yes. If No, need to understand why?
9 | Are your calibration gas flow rates set correctly to the analyzer in question?
10 | Are the calibration bottle regulator settings properly set?
11 | Is there only one component "Out of Control" or multiple components? |  |  | If only one component is "Out of Control", then you are probaly dealing with a component (analyzer) specific issue. If all measurements are "Out of Comtrol" there is a systemic problem that needs to be addressed?
12 | Does the analyzer "Repeat" on "Zero" gas? |  |  | If the analyzer(s) are not repeating on zero or span gas, the issue is related to the sample system or most likely the analyzer.
13 | Have you conducted a Bias Test? (Direct vs. Local Validation) |  |  | Perform a Bias (Line Loss Test) to insure a sub component is not causing the issue. The Bias Test should be conducted Quarterly.
14 | Did the Analyzer pass the Bias Test? |  |  | If No, you need to correct this before proceeding.
15 | Is the O2 reading on calibration gas above the established UCL? |  |  | If Yes, you may have an air leak in the system.
16 | Are the CO, NOx or SO2 reading on calibration gas below the establised LCL? |  |  | If Yes, you may have an air leak in the system.
17 | Have you conducted a Converter Efficiency Test? (Nox measurement only) |  |  | If No, then you need to conduct one. Must be done Quarterly.
18 | Are there any signs of water in the system? |  |  | If Yes, then you need to correct this problem immediately. Electric Chillers CANNOT handle water in the liquid phase PRIOR to the chiller. If the sample flow rate is too high, this may result in exceeding the cooling capacity of the chiller.
19 | Is the Sample Probe and Line at operating temperature?
20 | Is the Sample Chiller at the proper operating temperature?
21 | If the system uses a "Permapure" type dryer, is it operting properly?
22 | Is the Sample Pump running? |  |  | If No, is there water in the water slip sensor?
23 | Is high vacuum or low sample outlet pressure seen? |  |  | High vacuum indicates blockage in the system prior to the pump. Low sample outlet pressure could be a sign of an issue with the sample pump.
24 | Is the analyzer rack or area where the analyzers located air conditioned? |  |  | If No, you need to consider that the equipment might be too hot. If the SQC chart was developed during a time that the area was properly cooled, you could be seeing an issue related to the controlled environment.
25 | Is there a recirulation fan installed on the cabinet / enclosure? If so, is it working? |  |  | Concern is over temperature of the sample.
26 | Have you verified the local analyzer reading and the reading at the remote device (DCS or PLC) are in agreement?
27 | Did you correct or discuss this problem with the Plant Analyzer Specialist BEFORE proceeding with recalibration of the system?
REVISION 0 - OCTOBER 2012
©Smith Analytical, LLC. Can only be used by specified Clients of Smith Analytical. Use by others without permission from Smith Analytical, LLC or for Commerical Applications without prior consent is strictly prohibited.

## Continuous Analyzer
SMITH ANALYTICAL, LLC
PO BOX 278, SEABROOK, TEXAS 77586
281-667-6463
WWW.SMITHANALYTICAL.COM
Purpose of this document is to provide a checklist for Analyzer Professionals to insure all proper steps are taken BEFORE the Analyzer is calibrated. In the majority of the cases (>95%) where Statistical Quality Control Chart (SQC) are used and analyzers are shown to be "In Control", when the analyzer goes "Out of Control" something has changed and this needs to be addressed before calibration of the unit is allowed. Additional items can and should be added for your particular location. ALL ITEMS MUST BE REVIEWED BEFORE RECALIBRATING AN ANALYZER THAT WAS PREVIOUSLY IN CONTROL.
SQC "OUT OF CONTROL" CHECK LIST
ITEM | DESCRIPTION | OK - YES/NO | N/A | Comment
1 | Was the Analyzer "In Control" Before This Calibration? |  |  | If Yes, then you have a New Problem. If No, then NO assumptions regarding system performance should be made.
2 | Have all standard operating parameters been reviewed to insure everything is correct? |  |  | If No, this needs to be done BEFORE proceeding. Make NO assumptions. Check the Factory documentation to insure everything is correctly set.
3 | Are all switches in the proper setting for the analyzer? |  |  | Check range switches.
4 | Has anyone worked on the analyzer recently? |  |  | If Yes, you need to understand everything this individual did to the unit. They possibly introduced a problem.
5 | Was the Calibration Gas Bottle Changed?
6 | Is the calibration bottle pressure below 200 PSIG? |  |  | If Yes, you should consider changing the bottle and then running the validation test again.
7 | Is the analyzer repeating on calibration gas? |  |  | Here the interest is precision of the analysis even though the unit is out of control. If the analyzer is NOT repeating on calibration gas, you have a "Y" axis issue.
8 | Is the calibration sample being introduced to the analyzer in the same manner as done previously? |  |  | Answer should be Yes. If No, need to understand why?
9 | Are your calibration gas flow rates set correctly to the analyzer in question?
10 | Are the calibration bottle regulator settings properly set?
11 | Is there only one component "Out of Control" or multiple components? |  |  | If only one component is "Out of Control", then you are probaly dealing with a component (analyzer) specific issue. If all measurements are "Out of Comtrol" there is a systemic problem that needs to be addressed?
12 | Does the analyzer "Repeat" on "Zero" gas? |  |  | If the analyzer(s) are not repeating on zero or span gas, the issue is related to the sample system or most likely the analyzer.
13 | Is the O2 reading on calibration gas above the established UCL? |  |  | If Yes, you may have an air leak in the system.
14 | If the analyzer is measuring for something other than Oxygen (Hydrogen, CO, NOx or SO2) is the reading on calibration gas below the establised LCL? |  |  | If Yes, you may have a leak in the system or may be cross containing the sample with process sample.
15 | Are there any signs of water in the system? |  |  | If Yes, then you need to correct this problem immediately.
16 | Is the Sample Probe and Line at operating temperature?
17 | Is the sample regulator at the proper pressure setting?
18 | Is the Sample Chiller at the proper operating temperature?
19 | If the system uses a "Permapure" type dryer, is it operting properly?
20 | Is the Sample Pump running?
21 | Is high vacuum or low sample outlet pressure seen? |  |  | High vacuum indicates blockage in the system prior to the pump. Low sample outlet pressure could be a sign of an issue with the sample pump.
22 | Is the analyzer rack or area where the analyzers located air conditioned? |  |  | If No, you need to consider that the equipment might be too hot. If the SQC chart was developed during a time that the area was properly cooled, you could be seeing an issue related to the controlled environment.
23 | Is there a recirulation fan installed on the cabinet / enclosure? If so, is it working? |  |  | Concern is over temperature of the sample.
24 | Have you verified the local analyzer reading and the reading at the remote device (DCS or PLC) are in agreement?
25 | Did you correct or discuss this problem with the Plant Analyzer Specialist BEFORE proceeding with recalibration of the system?
REVISION 0 - OCTOBER 2012
©Smith Analytical, LLC. Can only be used by specified Clients of Smith Analytical. Use by others without permission from Smith Analytical, LLC or for Commerical Applications without prior consent is strictly prohibited.

## pH
SMITH ANALYTICAL, LLC
PO BOX 278, SEABROOK, TEXAS 77586
281-667-6463
WWW.SMITHANALYTICAL.COM
Purpose of this document is to provide a checklist for Analyzer Professionals to insure all proper steps are taken BEFORE the Analyzer is calibrated. In the majority of the cases (>95%) where Statistical Quality Control Chart (SQC) are used and analyzers are shown to be "In Control", when the analyzer goes "Out of Control" something has changed and this needs to be addressed before calibration of the unit is allowed. Additional items can and should be added for your particular location. ALL ITEMS MUST BE REVIEWED BEFORE RECALIBRATING AN ANALYZER THAT WAS PREVIOUSLY IN CONTROL.
SQC "OUT OF CONTROL" CHECK LIST
ITEM | DESCRIPTION | OK - YES/NO | N/A | Comment
1 | Was the Analyzer "In Control" Before This Calibration? |  |  | If Yes, then you have a New Problem. If No, then NO assumptions regarding system performance should be made.
2 | Have all standard operating parameters been reviewed to insure everything is correct? |  |  | If No, this needs to be done BEFORE proceeding. Make NO assumptions. Check the Factory documentation to insure everything is correctly set.
3 | Are all switches in the proper setting for the analyzer? |  |  | Check range switches.
4 | Has anyone worked on the analyzer recently? |  |  | If Yes, you need to understand everything this individual did to the unit. They possibly introduced a problem.
5 | Is the analyzer signal stable on the "Standard"? |  |  | If No, then stop and determine issue with the analyzer or related sub-assembly per the Manufacturers recommendations (i.e. pH Probe Simulator)
6 | Are you using "Fresh" pH Standards? Have they exceeded the "Expiration" date?
7 | Is the container you are using for the standards clean?
8 | Have you checked the pH probe to insure it is clean?
9 | Is the temperature compensator working?
10 | Did you correct or discuss this problem with the Plant Analyzer Specialist BEFORE proceeding with recalibration of the system?
REVISION 0 - OCTOBER 2012
©Smith Analytical, LLC. Can only be used by specified Clients of Smith Analytical. Use by others without permission from Smith Analytical, LLC or for Commerical Applications without prior consent is strictly prohibited.